Laboratory: _________________________________________ Date Implemented___________

Address: ___________________________________________

___________________________________________________

___________________________________________________

ProAdvantage® by NDC Strep A Test

- Test Principle

The Strep A Test is a qualitative, lateral flow immunoassay for the detection of Strep A carbohydrate antigen in a throat swab. In this test, antibody specific to Strep A carbohydrate antigen is coated on the test line region of the strip. During testing, the extracted throat swab specimen reacts with an antibody to Strep A that is coated onto particles. The mixture migrates up the membrane to react with the antibody to Strep A on the membrane and generate a red line in the test region. The presence of this red line in the test region indicates a positive result, while its absence indicates a negative result. To serve as a procedural control, a red line will always appear in the control region if the test has been performed properly. If a red control line does not appear, the test result is not valid.

- Specimen Collection/Treatment

| A. Specimen: | Throat swab |
| --- | --- |
| B: Specimen Container: | Clean, dry plastic tube |
| C. Specimen Storage: | 8 hours at room temperature in dry plastic tube or 72 hours at 2-8°C. |
| D. Handling Precautions: | Patient samples, controls, and test strips should be handled as though they could transmit disease. Observe established precautions against microbial hazards. |

- Reagents and Equipment

- Reagents and Materials Provided

| - Component | - Content | - Quantity |
| --- | --- | --- |
| - Test strips |  | - 25 |
| - Disposable extraction test tubes |  | - 25 |
| - Sterile swabs |  | - 25 |
| - Strep A Reagent A | - 2M Sodium nitrite | - 1 |
| - Strep A Reagent B | - 0.4M Acetic acid | - 1 |
| - Strep A Positive Control | - Non-viable Strep A; 0.09% Sodium azide | - 1 |
| - Strep A Negative Control | - Non-viable Strep C; 0.09% Sodium azide | - 1 |
| - Procedure Card |  | - 1 |
| - Package Insert |  | - 1 |
| - Workstation |  | - 1 |

- Materials not Provided

- Timer

- Storage and Stability

The kit can be stored at room temperature or refrigerated (2-30°C). The test strip is stable through the expiration date printed on the sealed pouch. The test strip must remain in the sealed pouch until use. DO NOT FREEZE. Do not use beyond the expiration date.

- Quality Control

Internal Quality Control

Internal procedural controls are included in the test. A red line appearing in the control region (C) is an internal positive procedural control. It confirms sufficient specimen volume and correct procedural technique. A clear background is an internal negative background control. If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test result.

External Quality Control

It is recommended that a positive and negative external control be evaluated to verify proper test performance with each new lot, each new shipment, monthly as a check on storage, each new untrained operator and as otherwise required by your lab internal quality system procedures. External positive and negative controls are supplied in the kit. Alternatively, other Group A and non-Group A Streptococcus ATCC reference strains may be used as external controls. Some commercial controls may contain interfering preservatives; therefore, other commercial controls are not recommended.

Procedure for External Quality Control Testing

- Add 4 full drops of Reagent A and 4 full drops of Reagent B into an extraction test tube. Tap the bottom of the tube gently to mix the liquid.

- Add 1 full drop of positive or negative control solution into the tube, holding the bottle upright.

- Place a clean swab into the tube. Rotate the swab 10 times in the tube. Leave the swab in the tube for 1 minute. Then press the swab against the side of the tube and squeeze the bottom of the tube while removing the swab so that most of the liquid stays in the tube. Discard the swab.

- Continue with Step 4 of Directions For Use.

## Precautions

- For professional in vitro diagnostic use only. Do not use after expiration date.

- Do not eat, drink or smoke in the area where the specimens and kits are handled.

- Handle all specimens as if they contain infectious agents. Observe established precautions against microbiological hazards throughout the procedure and follow the standard procedures for proper disposal of specimens.

- Wear protective clothing such as laboratory coats, disposable gloves and eye protection when specimens are assayed.

- WARNING: Reagent A is harmful if swallowed or adsorbed through skin. May cause eye irritation.

- CAUTION: Reagent B may cause skin, eye and respiratory tract irritation.

- The positive and negative controls contain sodium azide (NaN3) as a preservative.

- Do not interchange reagent bottle caps.

- Do not interchange external control solution bottle caps.

- Humidity and temperature can adversely affect results.

- Test Procedure

Allow the test strip, reagents, and/or controls to reach room temperature (15-30°C) prior to testing.

- Remove the test strip from the sealed foil pouch and use it as soon as possible. Best results will be obtained if the test is performed immediately after opening the foil pouch.

- Hold the Reagent A bottle upright and add 4 full drops (approximately 240 µL) to an extraction test tube. Reagent A is red in color. Hold the Reagent B bottle upright and add 4 full drops (approximately 160 µL) to the tube. Reagent B is colorless. The addition of Reagent B to Reagent A changes the color of the solution from red to pale yellow. Tap the bottom of the tube gently to mix the liquid.

- Immediately add the throat swab into the tube of pale yellow solution. Rotate the swab vigorously 10 times in the tube. Leave the swab in the tube for 1 minute. Then press the swab against the side of the tube and squeeze the bottom of the tube while removing the swab so that most of the liquid stays in the tube. Discard the swab.

- With arrows pointing down, place the test strip into the tube of solution and then start the timer. If the procedure is followed correctly, the liquid should be at or just below the maximum line (MAX) on the test strip. See the illustration below.

- Leave the strip in the tube and read the result at 5 minutes. The result is invalid after 10 minutes.

- Interpretation of Test Results

POSITIVE*: Two distinct red lines appear. One line should be in the control line region (C) and another line should be in the test line region (T). A positive result means that Strep A specific antibodies were detected in the sample.

*NOTE: The shade of the red color in the test line region (T) will vary based on the amount of Strep A IgG specific antibodies in the sample. Any shade of red in the test line region (T) should be considered positive.

NEGATIVE: One red line appears in the control line region (C). No apparent red or pink line appears in the test line region (T). A negative result means that Strep A IgG specific antibodies were not found in the sample or are below the detection limit of the test.

INVALID: No line appears in the control line region (C). If this occurs, read the directions again and repeat the test with a new test strip. If the result is still invalid, stop using the test kit and contact your distributor.

- Limitations

- The Strep A Test – Throat Swab is for in vitro diagnostic use only. The test should be used for the detection of Strep A antigen in throat swab specimens only. Neither the quantitative value nor the rate of increase in Strep A antigen concentration can be determined by this qualitative test.

- This test will only indicate the presence of Strep A antigen in the specimen from both viable and non-viable Group A Streptococcus bacteria.

- A negative result obtained from this kit should be confirmed by culture. A negative result may be obtained if the concentration of the Strep A antigen present in the throat swab is not adequate or is below the detectable level of the test.

- Excess blood or mucus on the swab specimen may interfere with test performance and may yield a false positive result. Avoid touching the tongue, cheeks, and teeth4 and any bleeding areas of the mouth with the swab when collecting specimens.

- As with all diagnostic tests, all results must be interpreted together with other clinical information available to the physician.

- Expected Values

Approximately 15% of pharyngitis in children ages 3 months to 5 years is caused by Group A beta-hemolytic Streptococcus.5 In school-aged children and adults, the incidence of Strep throat infection is about 40%.6 This disease usually occurs in the winter and early spring in temperate climates.3

- Performance Characteristics

Using three medical centers for evaluation, a total of 499 throat swabs were collected from patients exhibiting symptoms of pharyngitis. Each swab was rolled onto a sheep blood agar plate, and then tested by the Strep A Test – Throat Swab. The plates were further streaked for isolation, and then incubated at 37°C with 5-10% CO2 and a Bacitracin disk for 18-24 hours. The negative culture plates were incubated for an additional 18-24 hours. Possible GAS colonies were subcultured and confirmed with a commercially available latex agglutination grouping kit.

Of the 499 total specimens, 375 were found to be negative by culture and 124 were found to be positive by culture. During this study, two Strep F specimens yielded positive results with the Test. One of these specimens was re-cultured, then re-tested and yielded a negative result. Three additional different Strep F strains were cultured and tested for cross-reactivity and also yielded negative results.

|  |  | Culture |  |
| --- | --- | --- | --- |
|  |  | + | - |
| Strep A Test- Throat Swab | + | 120 | 20 |
|  | - | 4 | 355 |

Sensitivity:	120/124 = 97% (91%-99%)*

Specificity:	355/375 = 95% (92%-97%)*

Accuracy:	475/499 = 95% (93%-97%)*

Prevalence:	124/499 = 25%

PPV (+):	120/140 = 86% (79%-91%)*

NPV (-):	355/359 = 99% (97%-100%)*

* Denotes a 95% Confidence Interval

| Positive Culture Classification | Rapid Strip/Culture | % Correct |
| --- | --- | --- |
| Rare | 10/11 | 91% |
| 1+ | 9/9 | 100% |
| 2+ | 17/19 | 89% |
| 3+ | 36/37 | 97% |
| 4+ | 48/48 | 100% |

Cross-Reactivity

The following organisms were tested at 1.0 x 107 organisms per test and were all found to be negative when tested with the Strep A Test – Throat Swab. No mucoid-producing strains were tested.

| Group B Streptococcus | Group C Streptococcus |
| --- | --- |
| Group F Streptococcus | Group G Streptococcus |
| Streptococcus pneumoniae | Streptococcus sanguis |
| Streptococcus mutans | Enterococcus faecalis |
| Staphylococcus aureus | Staphylococcus epidermidis |
| Corynebacterium diphtheriae | Serratia marcescens |
| Candida albicans | Klebsiella pneumoniae |
| Pseudomonas aeruginosa | Bordetella pertussis |
| Neisseria meningitidis | Neisseria gonorrhoeae |
| Neisseria sicca | Neisseria subflava |
| Branhamella catarrhalis | Haemophilus influenza |

- POL Studies

Three physicians’ offices were used to conduct an evaluation of the Strep A Test – Throat Swab. Personnel with various educational backgrounds performed the testing. Each physician’s office tested a randomly coded panel of samples consisting of negative (20), low positive (20), and medium positive (20) for three days. The results obtained had a 96% correlation with the expected results.

- References

- Manual of Clinical Microbiology, 6th Edition, ASM Press, p. 299-307.

- Webb, KH. Pediatrics (Feb 1998), 101:2, 2.

- Bisno AL, Gerber MA, Gwaltney JM, Kaplan EL, Schwartz RH. Clinical Infectious Diseases (1997), 25, 574-83.

- Shea, Y.R., Specimen Collection and Transport, in Clinical Microbiology Procedures Handbook, Isenberg, H.D., American Society of Microbiology, Washington, D.C., 1.1.1-1.1.30, 1992.

- Nussinovitch, M, Finkelstein Y, Amir J, Varsano, I. Clinical Pediatrics (June 1999), 357-360.

- Woods WA, Carter CT, Stack M, Connors Jr AF, Schlager TA, Southern Medical Journal (May 1999), 491-492.

- ProAdvantage® by NDC Strep A Test Package Insert

Test Procedure Review

| Supervisor | Date Reviewed | Supervisor | Date Reviewed |
| --- | --- | --- | --- |