Laboratory: 						 		Date Implemented:

Address:

ProAdvantage® by NDC Infectious Mononucleosis Test Device (Whole Blood)
Laboratory Procedure

- Test Principle

- The Infectious Mononucleosis Test Device (Whole Blood) is a qualitative membrane strip based immunoassay for the detection of IM heterophile antibodies in whole blood, serum or plasma. In this test procedure, bovine erythrocyte extracted antigen is coated on the test line region of the device.  The sample reacts with bovine erythrocyte extracted antigen coated particles that have been applied to the label pad.  This mixture migrates chromatographically along the length of the test strip and interacts with the coated bovine erythrocyte extracted antigen. If the sample contains IM antibodies, a colored line will appear in the test line region indicating a positive result.  If the sample does not contain IM heterophile antibodies, a colored line will not appear in this region, indicating a negative result.  To serve as a procedural control, a colored line will always appear at the control line region, indicating that proper volume of specimen has been added and membrane wicking has occurred.

- Specimen Collection/Treatment

| A. Specimen: | Whole Blood specimen from venipuncture or fingerstick. |
| --- | --- |
| B. Specimen Collection: | Venipuncture Whole Blood samples: Collect anti-coagulated blood sample (sodium or potassium heparin, sodium or potassium EDTA, sodium or potassium citrate and sodium oxalate) following standard laboratory procedures. Fingerstick Whole Blood samples: • Wash the patient’s hand with soap and warm water or clean with an alcohol swab. Allow to dry. • Massage the hand without touching the puncture site by rubbing down the hand towards the fingertip of the middle or ring finger. • Puncture the skin with a sterile lancet. Wipe away the first sign of blood. • Gently rub the hand from wrist to palm to finger to form a rounded drop of blood over the puncture site. • Touch the end of the capillary tube to the blood until filled to the line; avoid air bubbles. • Place the bulb onto the top end of the capillary tube. • Squeeze the bulb to dispense the whole blood. |
| C. Specimen Storage: | Do not leave the samples at room temperature for prolonged periods. Whole blood collected by venipuncture should be stored at 2°-8°C (36°-46°F) if the test is to be run within 2 days of collection. Whole blood collected by fingerstick should be tested immediately. Do not freeze whole blood samples. Bring samples to room temperature prior to testing. If samples are to be shipped, they should be packed in compliance with federal regulations covering the transportation of etiologic agents. |
| D. Handling Precautions: | Handle all specimens as if they contain infectious agents. Observe established precautions against microbiological hazards throughout the procedure and follow the standard procedures for proper disposal of specimens. |

- Reagents and Equipment

- Reagents and Materials Provided

- 20 individually packaged test devices

- 20 disposable sample droppers

- 1 containers of disposable heparinized capillary tubes

- 1 dispensing bulbs

- 1 Mono sample buffer (5mL): containing 0.09% sodium azide

- 1 Mono negative control (1mL): Diluted human plasma, 0.09% sodium azide

- 1 Mono positive control (1mL): Diluted human plasma containing IM heterophile antibodies, 0.09% sodium azide

- 1 procedure card

- 1 directional inserts: 1 (CLIA waived on front and moderately complex on back)

- Materials Required but not Provided

- Sample collection container (for venipuncture whole blood)

- Lancet (for fingerstick whole blood only)

- Timer

- Storage and Stability

The kit can be stored at room temperature or refrigerated 2°-30°C (36°-86°F).  The test device is stable through the expiration date printed on the sealed pouch.  The test device must remain in the sealed pouch until use.  DO NOT FREEZE.  Do not use beyond the expiration date.

- Quality Control

Internal Quality Control

Internal procedural controls are included in the test.  A red line appearing in the control region (C) is an internal positive procedural control.  It confirms sufficient sample volume and correct procedural technique.  A clear background is an internal negative background control. If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test result.

External Quality Control

It is recommended that external positive and negative controls be tested with each new kit, lot or shipment of product, with each change in operator with in the test kit, weekly as a check on continued storage conditions, and as otherwise required by your laboratory’s internal quality system procedures.  External positive and negative controls are supplied in the kit.  If controls do not perform as expected, assay results are invalid.

Procedure for External Quality Control Testing

Using the positive or negative external controls in place of a patient sample, add 1 drop of positive or negative control solution to a sample well (S) of a new test device, then add 1 drop of Sample Buffer.  Start the timer.  Continue with Step 3 in the Test Procedure section.

Remedial Actions

- If unexpected results are seen when running the controls, review the Direction for Use, Interpretation of Results and Limitations sections and repeat the test with another device. If the result is still invalid, stop using the test kit and contact your distributor.

## Precautions

- For professional in vitro diagnostic use only.

- Do not use after expiration date.

- Do not eat, drink or smoke in the area where the specimen samples and kits are handled.

- Handle all specimens and controls as if they contain infectious agents.  Positive and negative controls contain human plasma.  Observe established precautions against microbiological hazards throughout all procedures and follow the standard procedures for proper disposal of specimen samples.

- The positive and negative controls contain sodium azide as a preservative, which may form potentially explosive metal azide if it reacts with lead or copper plumbing.  Large quantities of water should be used to flush discarded controls down a sink.

- Wear protective clothing such as laboratory coats, disposable gloves and eye protection when specimen samples are assayed.

- Humidity and temperature can adversely affect results.

- Do not reuse the test.

- Discard the test device if package is torn, ripped or if device itself is damaged.

- Do not mix buffer and controls from different lots.

- Test Procedure

Allow the test device, sample, buffer and controls to reach room temperature 15° - 30°(59°- 86°F) before testing.

- Remove the test device from the foil pouch and use it as soon as possible.  For best results, perform the test immediately after opening the foil pouch.

- Place the test device on a clean and level surface.

For Whole Blood (Venipuncture) samples:

- Hold the dropper upright and add 2 drops of whole blood (about 50 L) to the sample well (S) of the test device.

- Add 1 drop of Sample Buffer to the sample well.

- Start the timer.

For Whole Blood (Fingerstick) samples:

- Add one capillary tube of blood (about 50 L) to the sample well (S) of the test device.

- Add 1 drop of Sample Buffer to the sample well.

- Start the timer.

- Wait for the red line(s) to appear. The result should be read at 5 minutes. The background should be clear before the result is read.

NOTE: Low titers of IM heterophile antibodies might result in a weak line appearing in the test line region (T) after a long period of time.  Do not read the result after 10 minutes.

- Interpretation of Test Results

POSITIVE*: Two distinct red lines appear.  One line should be in the control line region (C) and another line should be in the test line region (T).  A positive result indicates that IM heterophile antibodies were detected in the sample.

* NOTE: The shade of the red color in the test line region (T) will vary based on the amount of IM heterophile antibodies in the sample.  Any shade of red in the test line region (T) should be considered positive.

NEGATIVE: One red line appears in the control line region (C).  No apparent red or pink line appears in the test line region (T).  A negative result means that IM heterophile antibodies were not found in the sample or are below the detection limit of the test.

INVALID: No line appears in the control line region (C).  If this occurs, read the directions again and repeat the test with a new test device.  If the result is still invalid, stop using the test kit and contact your distributor.

- Limitations

- The Infectious Mononucleosis Test Device (Whole Blood) is for in vitro diagnostic use only. The test should be used for the detection of IM heterophile antibodies in whole blood, serum and plasma samples only. Neither the quantitative value nor the rate of increase in Mononucleosis antibody concentration can be determined by this qualitative test.

- The Infectious Mononucleosis Test Device (Whole Blood) will only indicate the presence of IM heterophile antibodies in the sample and should not be used as the sole criteria for the diagnosis of Mononucleosis infection.

- Grossly hemolyzed samples will yield invalid results.  Strictly follow the Directional Insert instructions to obtain accurate results.

- As with all diagnostic tests, all results must be interpreted together with other clinical information available to the physician.

- This assay has not been established for patients under 18 years of age.

- Expected Values

Epstein-Barr virus infection during adolescence or young adulthood causes infectious mononucleosis 35% to 50% of the time.1,5

The incidence of EBV-associated infectious mononucleosis in the USA has been estimated at 45 per 100,000 and is highest in adolescent and young adults- about 2 out of 1,000.  No seasonal pattern of EBV infection exists. The incubation period is 10 to 60 days, though 7 to 14 days is common for children and adolescents.

- Performance Characteristics

A total of 611 clinical samples were tested by three independent sites in a clinical study.  Slide agglutination served as the reference method for the study.  Serum, plasma and whole blood were also collected for the detection of IM heterophile antibodies by the Infectious Mononucleosis Test Device.

Of the 611 clinical samples collected, 185 were considered positive and 426 clinical specimens were considered negative by slide agglutination method.  The results for each sample matrix are summarized below.

| serum | Slide agglutination |  |  | Positive Agreement = 72/72 > 99% (95%-100%)** Negative Agreement = 168/168 > 99% (98%-100%)** Overall Agreement = 240/240 > 99% (98%-100%)** |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 72 | 0 |  |
|  | – | 0 | 168 |  |

| plasma | Slide agglutination |  |  | Positive Agreement = 58/58 > 99% (94%-100%)** Negative Agreement = 181/182 > 99% (97%-99%)* Overall Agreement = 239/240 > 99% (98%-99%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 58 | 1 |  |
|  | – | 0 | 181 |  |

| whole blood | Slide agglutination |  |  | Positive Agreement = 50/55 = 91% (80%-97%)* Negative Agreement = 76/76 > 99% (95%-100%)** Overall Agreement = 126/131 = 96% (91%-99%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 50 | 0 |  |
|  | – | 5 | 76 |  |

| ALL SPECIMENS | Slide agglutination |  |  | Positive Agreement = 180/185 = 97% (94%-99%)* Negative Agreement = 425/426 > 99% (98%-99.99%)* Overall Agreement = 605/611 = 99% (98%-99%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 180 | 1 |  |
|  | – | 5 | 425 |  |

| * | Denotes 95% Confidence Interval |
| --- | --- |
| ** | Denotes 97.5% Confidence Interval |

In addition, the clinical samples were tested with a commercially available rapid diagnostic test kit. 611 serum, plasma and whole blood specimens were used to compare the Infectious Mononucleosis Test Device (Whole Blood) to a comparator test.  The results showed a >99% agreement between the two test kits.  The results for each sample matrix are summarized below.

| serum | Comparator test |  |  | Positive Agreement= 72/73 > 99% (93%-99%)* Negative Agreement= 167/167 > 99% (98%-100%)** Overall Agreement= 239/240 > 99% (98%-99%)** |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 72 | 0 |  |
|  | – | 1 | 167 |  |

| Plasma | Comparator test |  |  | Positive Agreement= 59/60 = 98% (91%-99%)* Negative Agreement= 180/180 > 99% (98%-100%)** Overall Agreement= 239/240 > 99% (98%-99%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 59 | 0 |  |
|  | – | 1 | 180 |  |

| Whole Blood | Comparator test |  |  | Positive Agreement= 50/51 = 98% (90%-99%)* Negative Agreement= 80/80 > 99% (96%-100%)** Overall Agreement= 130/131 > 99% (96%-99%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 50 | 0 |  |
|  | – | 1 | 80 |  |

| All Specimens | Comparator test |  |  | Positive Agreement= 181/184 = 98% (95%-99%)* Negative Agreement= 427/427 > 99% (99%-100%)** Overall Agreement= 608/611 > 99% (99%-99.9%)* |
| --- | --- | --- | --- | --- |
|  |  | + | – |  |
| Infectious Mononucleosis Test Device | + | 181 | 0 |  |
|  | – | 3 | 427 |  |

| * | Denotes 95% Confidence Interval |
| --- | --- |
| ** | Denotes 97.5% Confidence Interval |

Interfering Substances

No interference with the Infectious Mononucleosis Test Device (Whole Blood) results was observed in samples containing high levels of hemoglobin (up to 1,000 µg/dL), bilirubin (up to 1,000 mg/dL) and human serum albumin (up to 2,000 mg/dL).  The test results were also unaffected when the hematocrit was altered ranging from 20% to 60% and when icteric and lipemic samples were tested.

POL Studies

Three physicians’ offices were used to conduct an evaluation of the Infectious Mononucleosis Test Device.  Personnel with various educational backgrounds performed the testing.  Each physician’s office tested a randomly coded panel of samples consisting of negative (15), low positive (15), moderate positive (15) and invalid (15) for three days.  The results obtained had a >99% correlation with the expected results.

Non-Laboratory User Study

A total of 77 untrained, inexperienced, non-laboratory participants were enrolled at three separate locations to demonstrate that they could follow the product instructions and perform the Infectious Mononucleosis Test Device (Whole Blood) and obtain results similar to those obtained by trained laboratory technicians.  Each participant received four blinded spiked whole blood samples: one negative, one invalid, one low positive and one medium positive.

Study participants were instructed to follow the Instructional Insert and Procedure Card instructions to test the provided samples and record their test results.  No other instruction or training was given.  Upon completion of the test, participants filled out a brief questionnaire regarding the test procedure and ease of use of the labeling.  The following results were obtained:

| Site | Low Positive | Medium Positive | Negative | Invalid | Total Correct |
| --- | --- | --- | --- | --- | --- |
| A | 23/27=85% (66-96%)* | 25/27=93% (76-99%)* | 27/27>99% (87-100%)* | 27/27>99% (87-100%)* | 102/108=94% (88-98%)* |
| B | 25/27=93% (76-99%)* | 25/27=93% (76-99%)* | 27/27>99% (87-100%)* | 27/27>99% (87-100%)* | 104/108=96% (91-99%)* |
| C | 23/23>99% (85-100%)* | 23/23>99% (85-100%)* | 23/23>99% (85-100%)* | 23/23>99% (85-100%)* | 92/ 92>99% (96-100%)* |

*Denotes 95% Confidence Interval

- References

- Pediatr Clin North Am. 1997 Dec; 44(6):1541-56

- Omori, M. 2002 Mononucleosis. http://www.emedicine.com/EMERG/topic319.htm

- Linde, A. 1996, Scand J Infect Dis Suppl. 100:83-8

- Papesch, M. & Watkins, R. 2001 Clin. Otolaryngol. 26, 3-8

- CDC National Center for infectious Diseases: EBV & IM: www.cdc.gov/ncidod/diseases/ebv.htm

- ProAdvantage® by NDC Infectious Mononucleosis Test Device Package Insert

# Test Procedure Review

| Supervisor | Date Reviewed | Supervisor | Date Reviewed |
| --- | --- | --- | --- |